National Iranian Oil Company Safety, Health and Environment Management HSE Instructions for NIOC’s Chemical Labs Reviewing Translation; Confirmed by; Publication Date of English Version. Davood Naderi/ Laleh H. Abbasi Training & Performance Measurement Department; Mani Abdollahzadeh Rad/ HSE Manager; September 2018. Responsibility/Subject; Prepared by; Confirmed by; Approved by; Publication date. Name; Mostafa Sadeghpour/ Mahboubeh Keshavarz Head of Industrial Health Department/Senior Specialist of Industrial Health Department; Siavash Derafshi/ HSE Manager; Reza Khalili/ CEO' Deputy in Non Industrial Operation; Ali Kardor CEO; March 2017. Position. Revision No.: 0 Revision Date: March 21, 2017 Document Code: NIOC-HSE-OH-WI-031-00

Table of Contents

1. Introduction	6

2. Objective	6

3. Scope of Application	6

4. Approving and Implementing Authorities	6

4-1 Approving Authority	6

4-2 Implementing Authority	6

HSE Departments of NIOC’s subsidiaries and associate companies:	7

Personnel:	7

5. Definitions and Terms	7

5- 1 Chemicals:	7

5-2 Chemical Safety:	7

5-3 Chemicals Storage:	8

5-4 Chemicals Transport:	8

5-5 Use	8

5-6 Safety Data Sheet (SDS)	8

5-7 Globally Harmonized System Of Classification And Labeling of Chemicals (GHS)	8

6. Basics of Standard Lab Design	9

6-1 Safe Lab Environment (dimensions, walls and flooring, lighting, and windows)	13

6-1-1 Safe Lab Benching	14

6-1-1-1 Types of Benching	15

6-2 Lab Rooms	16

6-2-1 Types of Lab Rooms	16

2. Equipment and materials storage room	16

6-3 Safe Lab Sinks and Plumbing	17

6-3-1 Lab Sinks	17

6-3-2 Safe Plumbing Systems	18

6-3-3 Water and Wastewater Plumbing	18

6-4 Safe Lab Ventilation Systems	19

6-4-1 Environmental Requirements for the Installation and Operation of Ventilation Systems	20

6-4-2 Positioning of Ventilation Systems in Labs	20

6-4-3 Positioning of Hood Systems in Labs	20

6-4-4 Hood Operation Regulations	21

6-4-5-1 Chemical fume hoods	21

Types of fume hoods:	22

2. Bypass hoods	22

3. Auxiliary air hoods	24

6. Carcinogen fume hoods	24

8. Radioisotope fume hoods	25

9. Ductless fume hoods	25

6-4-5-2 Biosafety Hoods	25

Safety tips to consider when using biosafety hoods in the lab:	26

6-4-6 Lab Hood Safety Tips	26

6-5 Electric Power Systems in Labs	27

6-5-1 Hazardous Locations	27

The following points should be considered in the design of the lab’s electric power system:	28

6-5-2 Tips on Using Electrical Cables in Labs	30

6-5-3 Tips on using Electrical Devices in Labs	31

6-6 Lab Safety Cabinets	32

6-6-1 Types of Safety Cabinets	32

1. Ventilated safety cabinets	33

2. Acid/base corrosives storage cabinets	34

3. Flammable storage cabinets	35

6-6-2 Advantages of Safety Cabinets	37

6-7 Safety Shower and Eye Wash	37

6-8 Chemicals Storage	39

6-9 Standard Lab Chairs and Tables	41

7. General Lab Safety Tips	42

8- Chemicals Ordering, Receipt, and Storage	42

9. Emergency Planning	44

10. Chemicals Release and Spillage and Relevant Cleaning Methods	46

12. Personal Protection Equipment	47

13. Training	49

14. Differences between Safety Data Sheets (SDS) and Safety Labels	49

14-1 Proper Chemicals Labeling	51

15. Safe Use of Chemicals	51

16. Lab Inspection	52

17. Special Measures for Certain Hazards	52

18. Physical Damage Prevention	57

19. Waste Management in Labs	57

20. References	58

NIOC ‘s Waste Management Guideline - NIOC-HSE-EN-GU-006-00	58

HSE Department’s Waste Categorization and Coding Guideline - MOP-HSED-GI-302(0)	59

Appendix 1 - Danger Classification	60

## 1. Introduction

Due to the presence of different types of chemicals (gas, solid, liquid), including toxic, and potentially flammable and explosive substances, chemical Labs are considered potentially hazardous. Chemicals are found in rather smaller quantities in laboratories than in production processes. However, failure to observe the relevant safety and health regulations while handling such small quantities might lead to major lab accidents. Many lab accidents are the result of lack of knowledge of the dangers associated with such substances and non-observance of relevant safety regulations in chemical labs of industrial or academic (universities) centers and involve irreparable damage.

Thus, observing the relevant safety and health regulations in laboratories not only helps maintain the personnel's health but also prevents such accidents and their associated financial and reputation repercussions. Observing the regulations does not fully guarantee the prevention of such accidents. However, it can significantly reduce their severity and risk of occurrence.

## 2. Objective

Working at a chemical lab like working in other professions requires certain expertise, awareness, and competence. Due to the presence of hazardous chemicals in the lab and the direct exposure of lab staff to such substances, continuous and effective monitoring of these centers is both mandatory and imperative.

In the present document, the minimum safety, health, and environment requirements for chemical labs are introduced. These requirements should be taken into account in the design and operation of these centers. All subsidiaries and associate companies equipped with labs (regardless of the lab’s size) should follow this instruction and ensure it’s fulfillment. All companies shall strive to improve health and safety conditions i n their labs.

## 3. Scope of Application

The present instruction applies to all NIOC’s subsidiaries.

## 4. Approving and Implementing Authorities

### 4-1 Approving Authority

NIOC’s HSE Department shall be responsible for the approval and updating of the present instruction.

### 4-2 Implementing Authority

Lab supervisor at NIOC’s subsidiaries and associate companies:

- The lab supervisor shall be responsible for ensuring that all provisions stipulated in the present instruction are followed by the lab personnel, contractors and visitors.

- The activities of lab visitors should be monitored by the relevant lab supervisor.

- The lab supervisor shall be responsible for ensuring the safety of lab personnel and lab visitors during their activities, visits, and training programs. The lab supervisor should identify the potential hazards within the lab and work with other departments to eliminate such hazards.

### HSE Departments of NIOC’s subsidiaries and associate companies:

HSE departments of NIOC’s subsidiaries and associate companies shall be responsible for supervising the implementation of the provisions stipulated in the present document and their compliance with the internal requirements of their own companies, and development of safety and health programs and regulations, based on the relevant national projects and regulations, within their labs.

### Personnel:

All lab personnel should be aware of the provisions stipulated in the present instruction

and all the procedures related to their duties within the lab. Lab personnel should be aware of the provisions stipulated in the present instruction and all the procedures relevant to their duties. In addition, they should attend chemical labs safety and health training programs and inform their lab supervisor of any unsafe and dangerous condition within the lab.

## 5. Definitions and Terms

### 5- 1 Chemicals:

Chemicals are compounds of mineral or organic origin containing at least one chemical element. These naturally or synthetically occurring substances are used in various industries.

They exist in different forms, including gas, fume, liquid, and solid. They enter human body through different pathways, including ingestion, touch, and inhalation, causing various diseases.

The harmful effects of being in contact with these substances are determined by their type, entry point, quantity, and duration of contact.

### 5-2 Chemical Safety:

Prevention and management of short-term and long-term adverse effects of chemicals on human body and environment during production, application, and disposal process is known as chemical safety.

### 5-3 Chemicals Storage:

Chemicals storage is a place where chemicals and toxins of different forms (gas, liquid, and solid) are temporarily kept in accordance with the relevant requirements.

### 5-4 Chemicals Transport:

Chemicals transport refers to the transfer of chemicals packed or placed in a container from the manufacturing factory, to the seller, purchaser, and consumer through the land, sea, or air, or from the warehouse to the point where it will be used through a special process (using conveyor belts, pipes, pumps, etc.).

### 5-5 Use

Use refers to the use and application of chemicals in research and educational labs (or any other place) as primary ingredients, mediators, and/or supplements.

### 5-6 Safety Data Sheet (SDS)

SDS is a data sheet containing information on the potential hazards of chemicals and methods concerning their safe use and handling.

The SDS template proposed by International Labor Organization consists of 16 parts below:

1. Identification, 2. Hazard Identification, 3. Composition/information on ingredients, 4. First aid measures, 5. Firefighting, 6. Accident prevention, 7. Handling and storage, 8. Exposure controls/personal protection, 9. Physical and chemical properties, 10. Stability and reactivity, 11. Toxicological information, 12. Ecological information, 13. Disposal considerations, 14. Transport information, 15. Regulatory information, and 16. Other information.

### 5-7 Globally Harmonized System Of Classification And Labeling of Chemicals (GHS)

The GHS was approved in the 1992 Rio Conference on Environment and Development by the United Nations also called Earth Summit (1992).

## 6. Basics of Standard Lab Design

Design and construction of buildings of different applications should be based on the relevant occupational safety and health standards. The following are some of the organizations involved in the development of such standards and abbreviations referring to the standards, themselves:

- National Fire Protection Association (NFPA)

- Occupational Safety and Health Administration (OSHA)

- Building Officials and Code Administrators International (BOCA)

NFPA 45 Standard classifies labs into four groups, i.e., A. B. C, and D according to allowed quantities of flammable and combustible liquids. According to Figure 1, Class A and Class D labs are the most dangerous and safest, respectively.

| Maximum Quantities of Flammable and Combustible Liquids and Liquefied |  |
| --- | --- |
| Excluding Quantities in Storage Cabinets or Safety Cans | Including Quantities In Storage Cabinets or Safety Cans |

Figure 1. Classification of Labs based on the Allowed quantities of Flammable and Combustible Liquids

NFPA 45 Standard is known as the Lab Safety Standard. It is considered a valuable guideline used for the assessment of chemical labs. Building standards essentially deal with safe construction of buildings and fire protection and are less concerned with occupational health requirements. The following are the most important features of a lab designed based on the relevant safety standards proposed by OSHA and NFPA 45 Standard:

- Hoods should be installed in less frequented corners of the lab so that the lab could be easily evacuated in event of an accident.

- The lab should be well-ventilated so that the chemical pollutants are directed outdoors and all the lab personnel can enjoy a safe breathing air.

- The hood should be equipped with a velocity sensor to warn the lab personnel whenever the capture velocity is below the acceptable level.

- The safety shower and eye washing station should be located near the center of the lab so that that they are always just a few steps away (10 seconds).

- The lab personal should also be able to use them simultaneously.

- The space used as a chemicals storage should be as small as possible.

- The fire extinguisher should be regularly controlled and charged according

- to a planned schedule.

- The ceiling, walls, and floor of the lab and storage room should be washable. The floor should have sufficient slope to direct liquid spillage towards the drainage channels.

- The lab’s storage room should have as little doors and windows as possible. If necessary, use windows matte and high-resistant glasses.

- Climate and weather conditions should be considered while construction a lab building. Lab buildings should be made of fireproof and fire-retardant materials.

In designing a lab, all subsequent potential costs should be taken into account so that you can make the most of the space available. Beside the space needed for cabinets, columns, wall thickness, the staircase, the lift, and bathroom, an area of approximately 65% of the total room area should be taken as the empty space. The designers should take the lab personnel's wide range of needs, their comfort, energy efficiency, etc. into account. The lab working area and the associated rooms should have a minimum ceiling height of 3 m. The minimum usable space per person should be 12 m3. In labs with ceiling heights greater than 4 m, the minimum usable space per person is determined based on the ceiling height of 4 m.

Figure 2. A Standard Lab Design

In the above design, the area surrounding the bench is where flammable substances are kept. These substances should be as far from the exit door as possible.

The bench

area is separated from the rest of the lab's working area by a transparent barrier. When the lab's door is closed, the study area will be separated from the other areas. Thus, individuals working in the study room will not be consistently exposed to the pollutants released in the lab (The door surfaces should be smooth, non-absorbent, washable, and resistant to chemicals).

The atmospheric pressure difference between the lab area and the corridor is negative. Thus, the air at the area surrounding the bench should be as clean as the air flowing in the corridor. Transparent partitions make it possible for all lab personnel to know what is going on in each area of the lab. The entrance area is where the personnel can safely socialize or have a drink or snack. The door connecting the corridor and the lab area should not be opened outwards so that it does not hinder movement in the corridor (lab and storage doors should have separate keys and locks. Only the qualified personnel can enter the lab or the storage area. The network of facilities associated with supply of water, gas, and electricity to the lab and waste management should be designed to provide maximum safety. In addition, the facilities plan should be kept in an appropriate place at hand).

### 6-1 Safe Lab Environment (dimensions, walls and flooring, lighting, and windows)

A safe lab environment helps maintain lab personnel’s health and achieve the desired results. Temperature changes, non-observance of safety requirements, release of toxic fumes from different chemicals, etc. can surely make the lab an unsafe environment for the users.

Thus, before performing any test, make sure all the safe requirements are met and the lab’s personnel and the physical environment are protected against the potential hazards.

To ensure the safety of the lab’s physical environment, the following points should be taken into consideration:

- Labs should have standard dimensions so that sufficient space can be allocated for different activities.

- The minimum lab area, depending on the type of the lab, the required equipment, the experiments to be conducted, number of users, etc., is 78 m2. In some cases, if need be, the lab can have an area of 100 m2.

- The lab’s ceiling height is 3-3.5 m.

- The lab should have a minimum of two entry and exit doors with empty spaces up to a radius of 1.5 m.

- The window’s heights should not be less 100 cm. The total area of all the windows inside a lab should not excess one fifth it’s floor area.

- The lab should have adequate lighting (natural and artificial) especially in places where tests are conducted. This is of utmost importance in labs with potential microbial growth on dark and wet surfaces.

- Fluorescent lamps should be used to provide artificial mixed lighting. The lamps are installed at spacings of 1.5 m.

- Unprotected Ultraviolet (UV) light should not be used in the lab as it might cause discomfort to the users and disrupt their activities.

- The appropriate temperature and relative humidity for a standard lab are 25°C and 65%, respectively.

- The type of the lab and activities to be performed therein determine the features of its walls, ceiling, and flooring. The flooring should be resistant to both mineral and organic chemicals and be made of durable materials. It should also be replaceable if damaged. Today, vinyl floorings and tiles are considered a good option.

- The lab’s and storage’s walls and floor should be covered with tiles or other washable colored coverings as they are constantly exposed to toxic substances.

- The lab and storage's floor or the floor of other areas where lab users often frequent and substances are transferred should be smooth and flat and free from any holes, grooves, bumps, or other obstacles that might lead to people slipping or getting stuck.

- It should be noted that, other than bearing their own weight, walls will have to bear additional vertical loads.

- Thus, cabinets should not be mounted on walls with thicknesses less than 8 in (20.32 cm).

- The lab walls should be free from cracks and grooves and be fully connected to the ceiling and the floor so that organic substances and microorganisms cannot accumulate and grow in the potential gaps.

- The walls should be covered with durable materials. The wall covering should be washable and resistant to strong detergents and disinfectants.

- The lab walls should be washable to a height of at least 1.5 m and allow minimum water penetration.

- To create a brighter environment, walls should be painted with bright colors (reflection is among the properties of light and needs to be controlled).

- Use acid-resistant paints to protect walls against acid fumes. Durable and impermeable paints (e.g., epoxy paints), which are easily washed and cleaned, are recommended

- There could be need for stronger walls that are highly resistant to explosion. Firepro

- of walls limit the spread of fire and stretch from the ceiling to the floor without any gaps in between.

- The outer walls should also be resistant to the fire occurring inside the lab.

- The inner walls and ceiling can be painted white, light blue, light cream, o pale green.

### 6-1-1 Safe Lab Benching

Lab benches are prone to depreciation due to the presence of different types of chemicals within the lab environment and the wide range of physical and chemical test that are usually conducted therein. As a result, appropriate benches complying with the required technical and safety considerations should be used in the lab environment. The type of benches to be used is determined by the type of lab. Different benches have different features and properties. The following should be considered when selecting the best and safest benches for your lab:

- Beautiful design and choice of colors

- Appropriate size and dimensions in accordance with relevant lab standards Resistant to degradation, and chemical and physical agents

- Proper control and disposal of pollutants

### 6-1-1-1 Types of Benching

1. Double-sided or island benching

Double-sided benching is one of the most efficient benching designs used in labs. The design has a wide variety of advantages. For example, sinks can be installed on both sides of a double-sided bench, it has sufficient storage space for lab equipment and materials, and allows for the use of different power outlets.

2. One side or side benching

Single-sided benches have almost the same application as those of their double-sided counterparts. Lab devices are usually placed on single-sided benches.

3.Instrument benches

Instrument benches are a type of side bench. The back of instrument benches is covered with insulating plates to allow for the use of electrical connectors.

4. Sink benches

Sink benches are available in two types. The side sink bench is mounted on the wall.

The other type is attached to an island bench placed at the center of the lab. Sink benches are covered with uniform coatings made of anti-acid ceramics, granite or resin.

These uniform seamless coatings are beautiful and seamless, facilitating the cleaning process.

5. Balance tables

In labs, balance tables are used to prevent the transfer of ground vibrations to the balances, contributing to a more accurate measurement. Accordingly, the table’s top is made of stone, usually granite. The stone is placed on four rubber pads. Balance tables are equipped with a grounded power outlet mounted on a stand.

### 6-2 Lab Rooms

The lab’s pace is determined by the workload, the variety of tests to be performed therein, and the number of equipment and employees it is planned to accommodate. All labs, including research, operational, and service labs, should provide the lab staff with suitable working conditions so that they can be productive and achieve the desired results.

Thus, the provision of safe and appropriate working conditions should be taken as a priority when designing and equipping different labs, regardless of their size. In some special cases, for safety purposes, parts of the lab should be separated and enclosed (e.g., lab office, balance room, sensitive and special devices room). This will help prevent accidents in the lab.

The following section will introduce you to different types if lab rooms.

### 6-2-1 Types of Lab Rooms

1. Devices room

- Sensitive and/or expensive devices, such as atomic absorption spectrometers, radioactive devices, gas chromatography machines, which are not used on a daily basis, should be kept in a separate well-ventilated room with access to gas, water, and electricity.

- The number and type of equipment kept in the room has a significant impact on its design.

- Thus, factors, such as the device’s dimensions (length, width, and height), weight, voltage, and amperage and the necessary plumbing (as mentioned in the device catalog) should be taken into consideration.

- Rectifiers should be used in cases where the device in question operates with direct current (DC) power.

- Device manufacturers might recommend users to allocate a larger space to to their devices to help it function properly. Such recommendation should be closely followed. Easy access to the devices back and sides is necessary for maintenance, repair, and ventilation purposes.

### 2. Equipment and materials storage room

Every lab needs rooms for the storage of chemicals, and lab glassware, devices, and tools. A separate general storage room with shelving can also be used for storing such items. The location of the storage room should be carefully determined.

- Hazardous, flammable, volatile and toxic chemicals should be kept in a separate ventilated storage room.

- Producer instruction should also be followed when storing such substances.

- Storage rooms should be easily and instantly accessed.

- Storage rooms should well-ventilated and equipped with appropriate fire alarm and suppression systems.

- Safety requirements should be considered when determining the storage’s location, and selecting construction materials, air vents, hoods, and shelves.

### 6-3 Safe Lab Sinks and Plumbing

Proper disposal of chemicals and a provision of a safe and appropriate place for the washing of lab stuff are among the essential requirements of each lab.

Due to the inherent dangers resulting from the entry of chemical waste into the municipal sewage system and the need for sinks resistant to different types of chemicals, use of safe lab sinks and plumbing is necessary.

### 6-3-1 Lab Sinks

Safe lab sinks should have the following general features:

- Lab sinks should have an edge with adequate width to separate them from the worktop.

- Lab sinks should be equipped with cleanable sink drain catchers to avoid blockage and prevent solid stuff or broken glass entering the pipeline.

- Lab sinks should be easily cleaned and should not absorb chemicals.

- Lab sinks should be installed at an adequate height and have short enough faucets to prevent water spraying around when washing lab stuff.

- The most important feature of all lab sinks deals with their body as it must be made of materials resistant to chemicals and corrosive agents.

- Lab sinks can be categorized into the following groups by the material they are made of:

1- Epoxy resin sinks

2. Ceramic sinks

3. Stainless steel sinks

4. Polypropylene sinks

- The sinks to be installed in a lab are determined by the materials frequently used therein, lab dimension and the coatings covering the worktops.

The materials the sink is made of should be compatible with the coatings covering the lab’s worktops.

### 6-3-2 Safe Plumbing Systems

Plumbing should be done according to the available relevant standards and accurate and well thought out plumbing plans. The plumbing standards deal with pipe material, fittings, branch valves, and pipe installation methods. A safe lab plumbing system should have the following general features:

1. Resistant to corrosive substances and chemicals.

2. Long lifetime

3. High physical resistance

4. Easy installation and repair

5. Affordable

Metal pipes should never be used in lab plumbing systems due to their rather low resistance to organic acids. Despite their great resistance to corrosion, glass pipes should not be used in the lab plumbing system as they are highly fragile and rather expensive. Polypropylene pipes are the best option as they are resistant to a wide range of mineral and organic chemicals found in the lab’s sewage and are physically rigid. Furthermore, the lab plumbing system should prevent sewage gases enter the lab environment and avoid return flow. The vent plumbing system should be used for the plumbing of lab sinks. The vent plumbing system makes it possible to easily discharge water or wastewater from the sink into the sewage system and prevents sewer gases enter the sink and the lab environment. Essentially, the vent plumbing system maintains the pressure inside the sewer pipes at atmospheric level. Presence of water inside the gasket prevents gases entering the sink or the lab.

### 6-3-3 Water and Wastewater Plumbing

Polypropylene pipes are mainly used in the lab’s water and wastewater plumbing systems. Stainless steel pipes are used for gas piping.

Water and wastewater plumbing systems are of almost the same significance for labs. The only difference is that the flow of water in the pipeline is realized by the pressure applied by the water supply system, while wastewater flows through the pipeline due to the force of gravity. In addition, as mentioned earlier (“safety lab sinks and Plumbing”) sewer pipelines are equipped with siphons and vacuum breaker valves to avoid flow return and prevent hazardous foul-smelling gases enter the lab environment.

Vacuum breaker valves can be used to avoid flow return. These valves prevent negative pressure building in the pipeline and return of polluted water into the system. As shown in the figure below, vacuum breaker pumps have an air intake vent and a wastewater inlet valve.

Figure 3. Vacuum Breaker Valve Operation: (a) pipeline pressure is reduced; (b) water is flowing

The type of pipe to be used is determined by the operation conditions present in the lab. The pipes are categorized into the following groups based on the materials they are made of:

1. Carbon steel

2. Alloy steel

3. Stainless steel

4. Galvanized steel

Pipes, valves, and fittings should be selected and installed by taking into account the expert opinion of relevant specialists. The specialized considerations are not mentioned in the present document.

### 6-4 Safe Lab Ventilation Systems

Ventilation systems are installed in labs to provide a safe and conformable working place for those who deal with chemicals.

Fresh air should be constantly and mechanically introduced into the lab environment, while polluted air should be continuously evacuated through the hoods’ outlets.

Ventilation systems should be installed so that they prevent the return of polluted air to the lab environment/building.

### 6-4-1 Environmental Requirements for the Installation and Operation of Ventilation Systems

- Generally, the air flowing in the lab environment should be evacuated and replaced every 8 hours. However, it should be noted that the timing will vary by the type of activity performed in the lab and the level of pollution. For example, ventilation systems used in organic chemical labs prevent the release of toxic and hazardous fumes produced by chemicals or in chemical tubes into the lab environment. Thus, they are the first protective equipment to be considered for any lab.

- The fresh air should always enter the lab environment from the office. Thus, lab work should be performed in a separate area.

- Ventilation exhaust ducts, vents, and hoods should be installed in appropriate places to minimize air pollution.

- Ventilation systems used to deal with acid substances like perchloric acid should have a set of special features. For example, flexible air ducts and exhaust ducts should be acid-resistant, the system should be equipped with special filters and a water plumbing system for washing purposes.

### 6-4-2 Positioning of Ventilation Systems in Labs

- The lab’s general ventilation system should prevent the dispersion of pollutants in the lab.

- The system should be designed to combat eliminate pollution with minimum energy consumption.

- In the event of an increase in the lab’s heating load, the ventilation system should be adjusted to the new conditions.

- The ventilation system should be adjusted according to the lab staff’s attendance time in (weekends, night hours, etc.).

- When people are attending the lab: every 7 hours (minimum) (refer to standard ANSI/AIHA Z9.5)

- When no one is attending the lab: every 4 hours (minimum)

- The air introduced into the lab environment should come from the outdoors. The system should be periodically serviced to ensure is efficiency.

### 6-4-3 Positioning of Hood Systems in Labs

- Hoods should be placed at least 10 ft (3m) away from the lab doors (emergency doors are exceptions).

- The hoods should not interfere with movements inside the lab.

- Hoods should not be positioned so that air inlet vents are placed near or in front of each other (minimum distance of 2 m) as this might cause turbulence.

### 6-4-4 Hood Operation Regulations

- When using the system, make sure the lock keeping the glass sash closed is firmly secured and functional.

- Do not stick your head inside the hood when it’s turned on.

- The system’s functionality and efficiency should be annually assessed.

- If equipped with chemical sensors, make sure the pollution does return to the lab environment.

Refer to Standard ANSI/AIHA Z9.5 for more information on different types of hoods and ventilation systems and their operation within the lab environment.

6-4-5 Types of Lab Hoods

Hoods are categorized into two main groups: chemical

fume hoods and biosafety hoods.

### 6-4-5-1 Chemical fume hoods

Chemical fume hoods completely evacuate the chemical fumes and prevent their return into the lab environment. Back baffles should be adjusted according to the chemicals present in the environment as different chemical have different degrees of solubility and vapor density ratios.

It is recommended that fume hoods be equipped with alarms airflow indicators to ensure their safe and proper operation.

Fume hoods are basically equipped with two kind of alarms: Airflow alarm: This alarm is located at the upper left corner of the hood. The green light indicates that airflow is adequate. The amber light indicates that airflow is around the critical point. The red light indicates that the hood is not functioning properly.

Sash alarm: This alarm is located at the upper right corner of the hood. The alarm will turn on when the sash is pulled up.

It also allows the user to check the horizontal status of the sash.

Figure 4. Sash alarm (right) and airflow alarm (left)

### Types of fume hoods:

1. Conventional fume hoods

Conventional fume hoods basically operate the other types of hoods. The exhaust air at the front of the hood is pulled in by the fan and directed out of the building. It might also get filtered and returned back to the lab environment. By lowering the sash, the static pressure is reduced, while the airflow is increased. This will reduce the amount of air exiting the lab environment through the hood.

### 2. Bypass hoods

Bypass hoods are

designed to permit laboratory air pass from below and up the face sash opening. When the sash is lowered, it can produce a make-up airflow. The air drawn into the bypass flows through the hood. This will dilute the airflow and help filter out pollutants. Bypass hoods have two major advantage: (1) The airflow velocity near the work surface is constant. This will prevent airflow velocity changes that can interfere with the tests. (2) The static pressure is reduced.

### 3. Auxiliary air hoods

Auxiliary air hoods are equipped with an auxiliary air chamber eliminating the need for additional ventilation equipment required by other fume hoods. Thus, auxiliary air hoods are considered cost-effective. Nonetheless, these hoods have their own disadvantages too. A subtype of auxiliary air hoods evacuates the untreated air through the hood’s face sash opening. Thus, the user standing in front of the hood is exposed to untreated airflow.

Most experts recommend against the use of auxiliary air hoods. These hoods are hard to operate and maintain.

4. Walk-in hoods

Walk-in hoods are designed to be directly placed on the floor or a pad. They are at the same height as other devices.

As their height is not proportional to their width, walk-in hoods cannot produce the desired airflow and lack the advantages of their counterparts.

5. Self-contained hoods:

Self-contained hoods are usually used in cases where the lab needs a hood and the exhaust ducts are hardly available.

Commercial self-contained hoods are used in histology labs where substances such as formaldehyde, toluene, and alcohol are used. The newer ones allow for the use of optional filters when specific substances are used in the lab. Self-contained hoods draw in the room air through their filters.

The fan remains clean as the polluted air passes the filter before reaching the fan. Normally, it takes 1-2 years before the filter needs replacement. The major problem with self-contained hoods is that there is no warning alarm or indicator showing that the filter is filled. This problem has been resolved by using a substance that gives off a special odor when in contact with the filtered material. Self-contained hoods are more expensive that conventional fume hoods.

### 6. Carcinogen fume hoods

Working with carcinogens requires the use of high-quality hoods. These hoods protect users against any exposure to carcinogens.

7. Perchloric acid hoods Perchloric acid reacts violently with many substances. When evaporated inside the hood, it can generate explosive crystals. In order to evacuate these crystals and prevent their accumulation, perchloric acid hoods are made of stainless steel or polyvinyl chloride (PVC). The exhaust duct is equipped with water sprays and piping is done to supply the water needed for the internal water-wash system. Each perchloric acid hoods should have its own separate exhaust duct with all the joints and seams carefully welded.

The exhaust duct should have minimum twist and run the shortest path to the ceiling. The path must not be horizontal. It should have a 70-80 degree slope. The exhaust duct opening should be at least 10-15 ft. above the ceiling.

### 8. Radioisotope fume hoods

Radioisotope fume hoods are equipped with HEPA filters (High-Efficiency Particulate Air Filter) capable of absorbing 99.7% of particulates larger than 0.3 micron.

When using HEPA filters, the face airflow velocity should be controlled as the absorbed substances accumulated within the filter, reducing the airflow velocity. The HEPA filter will last longer if an affordable filter is placed before it.

Gaseous radioactive substances are not absorbed by HEPA filters. Thus, activated carbon or alumina filters are used to absorb such substances.

### 9. Ductless fume hoods

Ductless fume hoods pass the exhaust airflow through a filter or activated carbon and return it back to the lab environment. These hoods have limited application as a wide range of toxic chemicals are generally used in labs. Ductless fume hoods are only used in situations where no hazardous pollutant is present in the lab (in cases where the process can take place anywhere except under the hood).

Most ductless fume hoods are installed on top of lab caseworks. As these hoods are ductless, they should not be attached to any duct opening in the lab’s walls. Ductless fume hoods are largely made of stainless steel or fiberglass.

### 6-4-5-2 Biosafety Hoods

Studies and reports suggest that activities such as microbiological culture (microbial culture), acid preparation, and centrifugation produce air particulates in the air, causing people to become infected. Biosafety hoods are used to separate such particulates from the air we breathe. Biosafety hoods are classified as I, II (A and B), and III.

Each class provides a specific level of biosafety. Biosafety Hood Class I: These hoods are used to protect the lab environment and the individuals working therein. There are suitable for situations where the substances used in the lab pose a low or moderate risk.

Biosafety Hood Class II: These hoods are used to protect the environment, the individuals, the materials, and the products. The major difference between Class I and Class II is that the hoods belonging to the latter are equipped with HEPA filters to treat the room’s airflow.

Biosafety Hood Class III: These hoods provide the highest level of protection for the lab environment, personnel, and materials. They are used in cases where the substances used in the lab are extremely hazardous and full containment is required.

The HEPA filter is the most essential component of biosafety hoods belonging to this class.

### Safety tips to consider when using biosafety hoods in the lab:

- Turn on the hood at least 3-5 minutes before starting your work so that the air trapped within the hood is evacuated.

- Turn off the UV lamp and turn on the fluorescent lamp before starting your work. The reverse should be done after your work has finished.

- Use an appropriate disinfectant to clean the work surface, the internal surface of the sash, and the surfaces of all the materials, tools, and containers used within the hood.

- Ensure your hands and materials are appropriately placed within the cabinet to prevent airflow being pushed back.

- Ensure that the airflow is directed from the clean part to the contaminated part within the hood.

- Make sure the materials to be taken out of the hood are not contaminated.

- Make sure no object is placed in the upper part of the hood.

### 6-4-6 Lab Hood Safety Tips

- The exhaust fan should always be turned on.

- Baffles should be kept unobstructed.

- Large objects should not be placed on the work surface.

- Incompatible substances should not be placed next to each other.

- The emptier the hood, the higher its efficiency. Thus, never turn a hood into a storage.

- Chemical substances spilled within the hood should be cleaned.

- Making Sudden movements in front of the hood is prohibited.

- The performance certification label should be attached to the hood. The hood should have been tested within the past 12 months with its performance being certified.

- The alarm light should be tested. The green light should be on.

- Otherwise, the hood might have gone into the “Energy Saving Mode”, reducing its discharge.

- The objects should be kept at the back of the hood so far as possible, not less than 15-20 cm away from the air foil.

- Items put inside the hood should be placed 1 cm higher that the work surface so that the airflow can pass beneath them.

You should never assume that the hood is functioning properly all the time. All indicators should be checked.

Use a small piece of paper to check if the hood is functioning properly. If the it is not functioning properly, do not use hazardous and toxic chemicals inside the hood.

- Never fully lower the sash before making a bypass.

- Hold the door open at least 5 cm, especially when flammable substances are used in the hood.

- The door should be adjusted so that the operator is protected against any potential spillage.

- All operations should be performed at a minimum distance of 15 cm from the operator and inside the hood.

- The suction power might not be at its 100% at the front parts.

- In chemical fume hoods,

- the fan should by of type Ex.

### 6-5 Electric Power Systems in Labs

Individuals working in the lab are usually exposed to electrical hazards, including electric shock, fires, and explosions. In a lab, electrical hazards are caused by:

1. Faulty wiring, electrical devices, outlets and extension cords

2. Use of Unsafe procedures and methods

### 6-5-1 Hazardous Locations

Hazardous locations are classified according to the type and electrical conductivity level of materials used in their construction, and presence of flammable fumes and explosive dust.

National Electrical Code - NFPA 70 Standard divides hazardous location into three classes: I, II, and III.

1. Class I: It includes locations where flammable

gases and fumes are present in the air.

2. Class II: It includes locations where flammable or electrically conductive dust is present (dust with resistivity of greater than 10 ohm-cm is not electrically conductive).

3. Class III: It includes locations where flammable fibers are present.

Each class has 2 divisions. In Division 1, the hazardous conditions are the result of normal activities, frequent repair and maintenance operations, breakdown of equipment leading to the release of hazardous fumes, or falling of electrical devices. Division 2 includes hazardous locations where hazardous substances or the processes mentioned in Division 1 are present.

In Division 2, ventilation systems are used to keep the density of hazardous gases and particulates as low as possible.

Division 1 consists of groups A, B, C, and D. This categorization is based on the flammability level of the substances found in air.

Group A includes substances with the flammability limit ranging between 2.5% and 81% (e.g., acetylene).

Group B includes substances with

the flammability limit ranging between 4% and 75% and the maximum flash point of 37.8°C (e.g., ethylene oxide and hydrogen).

Group C includes substances with the flammability limit ranging between 2% and 30% (e.g., Acetaldehyde).

Group D includes substances with the flammability limit ranging between 1% and 17% (e.g., acetone).

### The following points should be considered in the design of the lab’s electric power system:

- An proper distribution of sufficient power outlets should be provided for the electrical equipment used in the lab, eliminating the need to use a variety of extension cords.

- Many older electrical equipment are designed based on two-phase circuit diagrams instead of their three-phase counterparts (phase, null, earth).

- Unfortunately, older circuits fail to provide proper connectivity and should be replaced with the newer ones as soon as possible.

- The electrical power systems should be spark-proof and equipped with an earthing system.

- The earthing system aims at preventing or, at least, minimizing electrical hazards including electric shock and fires.

- Fuses and circuit breakers are among the simplest devices used to stop excess currents.

- Circuit breakers can be re-used. However, fusses should be replaced by new ones.

- Fuses and circuit breakers should not be installed at places where flammable and explosive substances are present.

- Labs containing excess current or

- voltage sources should be equipped with a server.

- All wirings, connections, and power switches should be designed to withstand maximum current and voltage.

- Circuits should be protected by circuit breakers designed to stop currents exceeding their maximum limit.

- The connections should be enclosed with panels made of metal, plastic, etc.

- The active components of the panel breaker should be fully covered.

- If the panel breaker is placed in an isolated environment, those working at the lab should have access to it.

- The electrical service panel should not be used as a storage place for lab tools and equipment.

- Access to electrical panels, switches, and other electrical devices.

- Electrical equipment and circuits used in a lab should not be kept in places where there is risk of humidity exposure. In addition, they should not be used in places where the user is exposed to humidity.

- In places where flammable and explosive gases are probable to be found, use of spark-proof electrical devices is mandatory.

### 6-5-2 Tips on Using Electrical Cables in Labs

- Electrical wiring should be as simple as possible and easy to follow. A wiring diagram should also be drawn.

- Wire insulation and coating is needed to prevent electric shock. The coatings should be resistant to corrosive vapors, solvents, and high temperatures.

- Fragile coatings might crack and create hazardous conductive paths due to the presence of humidity in the lab.

- Extension cords are only allowed for temporary use. All safety precautions should be followed when using extension cords. They should not be placed in areas where individuals often frequent.

- Cracks resulting from use of bad materials, age, and other environmental factors increase the risk of short circuits.

- Extension cords with damaged coatings should be discarded.

Figure 5. Damaged Extension Cord

- Electrical cables might get damaged by the sharp edges of windows, doors, or iron cable ties or become worn out.

- Thus, cables should be protected against pressure and cutting. Electrical cables should not be placed in places where individuals often frequent.

- In places like the edges of windows an doors where damage to the coatings is likely, a physical cover like the external metal covers should be used to protect cables (Note: Many plastic coatings will soften when exposed to heat. Any contact between these cables and the external metal cover will increase the risk of electric shock. Thus, non-metal covers are highly recommended.).

### 6-5-3 Tips on using Electrical Devices in Labs

- All mobile and stationary devices should be properly equipped with an earthing system.

- Electrical equipment should be kept away from water and disconnected from the power source if not in use.

- Active electrical components should be protected against voltages of 50V or higher and accidental collisions.

- High-voltage equipment should be identified by warning signs.

- Damaged equipment should be marked with warning signs.

Figure 6. Warning Sign

- Electrical equipment used in the lab (e.g., mixer and oven) should not be placed near substances that might produce flammable fumes.

- Electrical equipment should only be repaired by experienced trained personnel.

- The repair tools should be kept in good conditions.

- Insulated protective gears should always be used when dealing with high

voltages.

- Safety cautions should be taken seriously when dealing with devices equipped with an earthing system.

### 6-6 Lab Safety Cabinets

When storing chemicals, storage conditions, arrangement methods, type of containers, etc. should be taken into consideration. The following points should be considered when storing lab materials on shelves and in ordinary boxes:

1. Cabinets should not be installed in places where they might interfere with the daily activities of the lab and prevent individuals to move freely in the lab environment.

2. Shelves should be able to bear the weight of chemicals. 3. All containers kept in the cabinet should be clearly and properly labeled. GHS (Globally Harmonized System of Classification And Labeling Of Chemicals) rules and regulations should be followed when labeling chemicals.

4. Incompatible substances should be separately stored. Improper storage of incompatible substances might lead to accidents (refer to the list of incompatible devices provided in Appendix 2).

5. Large quantities of hazardous substances should not be stored in cabinets. The excess quantities should be kept outside the lab environment.

6. Substances should be orderly stored on shelves and in special cabinets.

7. Keep in mind the chemicals’ expiration date and the date when you have first opened their container.

8. Chemicals should not be stored under the sink as they are likely to leak and react when exposed to humidity.

9. Fragile containers, especially those used for storing liquids, should be placed lower that the eye level.

10. Shelves should not be installed at heights where it would be hard to pick to reach. Lightweight chemicals and substances rarely used should be kept on the upper

shelves.

Refer to Standard ANSI/AIHA Z9.5 for more information on different types of safety cabinets (used for storage of chemicals) and their operation within the lab environment.

### 6-6-1 Types of Safety Cabinets

Besides ordinary cabinets, specialty cabinets can also be found in labs. These cabinets are used for the safe storage of chemicals and gas cylinders and are known as chemical storage cabinets or solvent cabinets.

These specialty cabinets are generally categorized into the following groups:

1. Ventilated safety cabinets

2. Acid/base corrosives storage cabinets

3. Flammable storage cabinets (solvent cabinets)

4. Gas cylinder cabinets (containing hazardous gases such as hydrogen, carbon monoxide, silane, and acetylene)

Some safety cabinets are dual-purpose. These dual-purpose safety cabinets are used to store acid/base corrosive substances along with the flammable chemicals. In the following sections, each category is

briefly described:

### 1. Ventilated safety cabinets

Volatile, corrosive, acid, base, and flammable liquids and solvents used in labs and different industries give off fumes that are harmful to our health and environment and have foul and unpleasant odors. Thus, such substances should be safely kept in special places to protect users and environment against their harmful effects. This will also help extend lifespan of the tools and equipment used in the lab. This issue is of special importance in roofed spaces with no ventilation or air-treatment system available, where toxic gases cannot be evacuated from the lab environment through hoods.

These cabinets are equipped with replaceable filters made of activated carbon.

Activated carbon is a porous substance made of carbon. It traps all the particulates and odors produced by chemicals into its numerous pores. Thus, in cabinets equipped with a fan and a carbon filter, the fan directs the air towards the activated carbon filter. The air passes through the filter and comes out safe, odorless, and breathable. Thus, these cabinets will solve the issue associated with the storage of chemicals with foul odors and harmful fumes. When kept in ordinary cabinets, the foul odors and harmful fumes will be present in the lab environment. Whenever an individual opens the door of an ordinary cabinet, he/she would inhale a considerable amount of toxic fumes which are substantially harmful for his/her lungs. Such fumes will also damage the tools and equipment used in the lab. However, if cabinets equipped with fans and activated carbon filters are used, the lab’s air will always remain safe. In addition, whenever you open the cabinet’s door, there will be no foul odor or toxic gases that can enter the lab environment (the ventilation system inside the cabinet is continuously working and the air is getting treated all the time).

Figure 7. Ventilated Safety Cabinet /Activated Carbon/Fan equipped with an Activated Carbon Filter

### 2. Acid/base corrosives storage cabinets

Acid and base liquids are among the substances that need to be carefully handles and stored. These liquids have a foul and unpleasant odor. When transferred out of the cabinet, there is always the risk that they spill over the cabinet or the user’s body parts.

The cabinet’s body should be resistant to acid and base liquids. They should have adjustable shelves and drawers with standard edges. They should also be easy to clean up as the glassware containing the liquids might fall and break inside the cabinet. These cabinets usually come with a special tray that prevents any leakage or spillage.

These cabinets are often made of anti-resistant materials, such as Polypropylene, HDPE (High-Density Polyethylene), and wood. Wooden cabinets should not be used to store oxidants such as nitric acid and pechloric acid. Metal cabinets are used to store oxidants; however, they will also sustain corrosion over time.

A solvent resistant tray

Figure 8. Corrosives storage cabinets

### 3. Flammable storage cabinets

Flammable substances are always challenging to store as there is often the risk of fires.

These cabinets prevent pyrophoric fumes from reaching flammable substances. When there is a fire outside the cabinet, flammable substances stored in these cabinets will be safe and there is no risk of explosion. When fire occurs outside in the lab environment, the area inside the cabinet should be isolated from the outside for 30 (min) to 90 (max) minutes (in 1000°C) so that the dire does not spread to the substances stored in the cabinet before the fire gets extinguished.

Most flammable storage cabinets are made of metals and have the following features:

1. The cabinet’s floor, ceiling, and walls should be made of double-layered iron panels. There should be a 1.5 cm spacing between the layers.

2.. Joints should be secured by welding, clinching, and other effective methods.

3. The cabinet door should be equipped with a three-point locking system.

4. A clearly visible “Flammable, Keep Fire Away” label should be attached to the cabinet.

Figure 9. Flammable storage cabinets

Flammable substances should be labeled based on the table below:

| Type of substance | Group | Flash Point | ‌Boiling Point |
| --- | --- | --- | --- |
| Flammable | IA | <73 °F | <100 °F |
| Flammable | IB | <73 °F | ≥ 100 °F |
| Flammable | IC | 73°F - < 100 °F |  |
| Flammable | II | 100 °F - <140 °F |  |
| Flammable | IIIA | 140 °F - < 200 °F |  |
| Flammable | IIIB | ≥ 200 °F |  |

Figure 10. Flammable substances categorization

4. Gas cylinder cabinets: This also applies to the storage of gas cylinders is enclosed areas. In other words, for safety purposes, cylinders should be kept and stored in special safety cabinets. This will eliminate the dangers associated with the storage of cylinders containing hazardous gases (e.g., leakage, falling, explosion, and fires), isolate them from the outside environment, and prevent the spread of potential fires occurring in the lab environment to the cylinders. If fumes originating from the containers of acid, base, flammable, and explosive substances are likely to get inside the cabinet, it should be equipped with a special ventilation system.

Figure 11. Gas cylinder Cabinet

### 6-6-2 Advantages of Safety Cabinets

1. Keeping the lab organized and in order.

2. Minimizing the hazards associated with the transfer of highly dangerous chemicals by reducing the need to transfer them from the storage point to the point to the consumption point

3. Creating a working environment with maximum safety

4. Preventing fire from entering the cabinet for a maximum period of 90 minutes and, in turn, keeping fire from spreading to the other parts of the building

5. Evacuating toxic and foul-smelling fumes through advanced ventilation systems

6. Preventing explosions

7. Extending the safe escape time needed during fires

8. Extending the time available for providing relief services and controlling and extinguishing fires

### 6-7 Safety Shower and Eye Wash

A) Safety Shower

- Safety showers should be installed in all areas where corrosive or toxic substances are used and stored. They should be accessible by lab personnel within at least 10 seconds.

- Safety showers should be easily accessible with no barrier or obstacle blocking the way.

- Safety showers should be installed in places where accumulation of water is unlikely. They should also be easy to clean.

- Make sure all the personnel are aware of the locations of safety showers within the lab environment and the neighboring areas.

- The safety shower location should be marked with green and white signs.

- Safety showers should be annually tested with the results recorded on specials labels or cards.

- The tank supplying water to the safety shower should supply at least 20 gallons (80 liter) of water per 15 minutes.

- The water pressure should not be too high to create discomfort and pain for the user when sprayed.

- Water temperature should be around body temperature (32°C - 35°C).

- Instructions on how to activate and deactivate the safety shower should be available near the shower and clearly visible.

- Safety showers could be activated by a pedal (by foot) or a handle (by hand).

- Safety showers should be inspected at least once per year.

- Safety showers should not be places bear electricity sources as the potential contact between water and electricity source poses the risk of electric shock.

B) Eye Wash

- The emergency eye wash should be used wherever corrosive substances, hot liquids or other eye irritants are used.

- Individuals should be aware of the locations of the eye washes and know how to operate them.

- Locations of emergency eye washes should be marked with special signs. These signs are usually white and green.

- Eye wash tests should be conducted once per week.

- Eye washes should be installed in the lab’s main area and be easily accessible.

- Eye washes should not be enclosed by physical barriers or be inly accessible through winding routes.

- An ideal eye wash station should be centrally located along the exit path or in any other appropriate place easily accessible by all personnel.

- It is recommended that the eye wash station be installed near where strong acids and bases are being frequently used.

- Small bottles of water (0.5 or 1 liter) are not an acceptable alternative option to eye washes. However, they can be helpful if used along the eye wash. In cases where both eyes have been in contact with and injured by chemicals, they should be unstoppably washed with a flow of potable water for 15-20 minutes.

- When using small water bottles (0.5 or 1 liter), make sure they are clean and uncontaminated.

- Eye wash stations should supply sufficient amount of water (min. 0.4 gallon (1.5 liter) per minute and pressure of approximately 30 psig).

- The valves used to turn on and off the water flow should be manually operated with minimum effort. However, there are several mechanisms to control the water flow. The most suitable mechanism involves the use of a pedal. The injured can push the pedal with his/her foot (eye washes cannot be manually activated/deactivated).

- It is recommended that eye washes be installed in the same place where safety showers are placed. Eye wash nozzles should be installed at a 6 inch distance from the wall and 33-45 inches above the floor.

- The water temperature should be around body temperature.

- All eye washes should be periodically inspected according to the relevant schedule. Any defect should be immediately repaired.

- A summarized version of the instructions should be available so the users know how to use the device, help the injured, and fully open one’s eye so that water could reach the damaged tissues.

- The lab personnel should receive continuous training on how to use the eye wash and react in situation when they need to use the device.

### 6-8 Chemicals Storage

According to the requirements stipulated in OSHA and NFPA 231 - General Storage, the chemicals storage should have the following features:

1. The storage rooms should be insulated against fires of different degrees of intensity. Doors should automatically close when fires occur. The Windows should be open. Overall, the standards set by the NFPA regarding the storage's doors and windows should be completely followed.

2. The following measures should be taken to prevent liquid leaking to other storage rooms: A) Making a slope at the entrance of the storage room with a minimum height of 10.16 cm; B) Making a groove in the storage’s floor with a minimum depth of 10.16 cm; and C) Making a furrow or a groove on one side of the storage room so that the leaked liquid accumulated therein and could not reach other chemicals kept in the room.

3. All storage rooms should be properly ventilated. According to NFPA standards (NFPA 231 - General Storage), different ventilation systems should be used for different storage rooms.

For example, in storage rooms where the ceiling height is greater than 3. m, two ventilation systems are needed.

4. Smoking and making fires in the storage rooms used for keeping flammable and combustible substances is prohibited. The “Flammable, No smoking” sign should be attached to the door of those storage rooms containing flammable substances.

5. The following points should be considered when keeping chemicals in the storage rooms:

- Each chemical should have an exclusive place prepared for its storage and should be returned to that place after being used in the lab

- All the chemicals and relevant equipment should be kept in the storage cabinets and on the shelves specifically designed for the storage room.

- Some barriers should be put in front of the containers to prevent them from falling on the ground.

- Chemical should not be stockpiled on the tables or above the cabinets.

- The maximum height of containers to be stored on the shelves is 1 m.

- All containers should be appropriately labeled with the receipt date written on the label for inventory control purposes.

- Toxic and foul-smelling substances should be kept in ventilated cabinets. Substances that need not be stored in ventilated cabinets should be kept in locked cabinets or on shelves having a front edge.

- Flammable liquids should be stored in special certified cabinets.

- Chemicals should not be stored in places where there are exposed to heat and direct sunlight.

- Chemicals should be stored based on their compatibility groups.

- A certain individual should be assigned at the storekeeper. The storekeeper should have a replacement for certain cases. The storage should be periodically inspected.

6. Attention to compatibility groups can be substantially helpful in managing the storage.

7. The following points should be considered in supplying fire extinguishing equipment needed at the chemicals storage.

- The equipment should be provided in sufficient numbers and amounts.

- The equipment should be kept at appropriate places.

- Both Manual and automatic operation of the equipment should be available.

- The equipment should be easily accessible.

- The equipment should not have reached their expiration date.

- The equipment should be periodically inspected according to the relevant schedules to ensure they operate flawlessly.

### 6-9 Standard Lab Chairs and Tables

The following points should be considered when selecting lab chairs:

1. Lab chairs should be designed according to the ergonomic principles to prevent musculoskeletal disorders in the long term.

2. The height chairs used in the lab office, and their seats and armrests should be standard and adjustable so that individuals with different body structures can easily use them.

3. The seats should be coated with polyurethane foams or vinyl, or be made of stainless steel as these materials do not absorb chemicals, eliminating the risk of fires. Use of fabric coatings is prohibited.

4. Chairs should be made of materials resistant to corrosive materials, solvents, and other chemicals. They should also be easy to clean.

5. Chairs without armrests should be used in the lab environment so far as possible as the armrests will limit the user’s movement.

6. When using a wheeled chair, make sure the chair can move swiftly on all types of surfaces.

7. Wheeled chairs should be equipped with a lock. The user can fix the chair in its place by pressing the lock.

8. It is recommended that chairs be equipped with a foot ring.

9. Chairs should not obstruct the exit and prevent individuals reaching fire extinguishers.

The following points should be considered when using lab tables:

Lab tables are mainly made of stainless steel or phenolic resin. These tables are usually wheeled and are easy to move around. Lab tables can be used at two heights. One height is suitable when the user is seated. The table is adjusted to the alternative height when the user is doing his/her work in a standing position.

1. The tables should be designed according to the ergonomic principles and have an appropriate height.

2. Tables should be made of materials resistant to chemicals and be easy to clean.

3. Lab tables should be placed near exit doors.

4. Tables should not be places near chemical fume hoods or their exhaust ducts.

9. Tables should not obstruct the exit and prevent individuals reaching fire extinguishers.

Note: Tables should be placed in office rooms or specific areas separated from the lab environment by walls.

## 7. General Lab Safety Tips

- Make sure the all access pints and entrances to the lab are under control (doors should be locked when no personnel in present in the lab).

- Unauthorized personnel should be requested to leave the lab.

- Fire-proof doors should be kept closed.

- Lab personnel are not authorized to work alone in the lab. In emergency cases, they are required to make the necessary arrangements so that the lab supervisor regularly checks on them while working alone.

- While working in the lab, do not touch your eyes, mouth, or skin lesions with your hands. Lab equipment should also not come into contact with your mouth, eyes, and any skin lesion.

- Food, drinks, smokes, chewing gums, medicine, and cosmetics are forbidden in places where chemicals or radiologic substances are potentially used or stored.

- The workplace should be kept clean. Lab corridors should be clean and unobstructed.

- Headphones should not be used for listening to music or talking with one another.

- Long hairs and short uniforms are forbidden in the lab.

- Ensure lab personnel have access to soap (preferably soft soap).

- Hands and other body parts should be washed with water and soap after working with chemicals or radiologic substances and before leaving the lab.

## 8- Chemicals Ordering, Receipt, and Storage

The following points should be considered when purchasing and storing chemicals in the lab:

1. Before purchasing chemicals, evaluate the supplier to see if he//she has the required qualification and capabilities.

2. The quantities and properties of chemicals to be delivered should correspond with the quantities and properties specified in the order.

3. When purchasing chemicals, ask the supplier or manufacture for the 16-item MSDS or SDS.

4. Given the requirement for asking the supplier for the chemical’s SDS or MSDS, the following points should be taken into consideration:

The wordings of the SDS should be comprehensible for the lab personnel. A copy of the SDS or its electronic version should be available for the lab personnel.

- Hazardous chemicals should not be used when no SDS is available.

- New SDSs should immediately be prepared and made available to the lab personnel.

- Not more than three years should have passed since the preparation of the SDS. Chemicals should be regularly investigated and the information with potential impacts on their safe handling should be updated in the relevant SDS.

- All the tools and equipment mentioned in the SDS (e.g., personal protection equipment, fire extinguishers, eye washes, and emergency safety showers) should be made available at the lab.

5. All the purchased chemicals should be marked with the special labels and warning signs by the supplier. The labels and signs should be used according to the relevant GHS standards.

6. The initial permits and authorizations should be obtained before placing any order for the chemicals.

7. Only order chemicals that can be safely stored and used with the methods and equipment available in the lab.

8. Order the chemicals in the minimum amount required.

9. Make the necessary arrangement for the receipt of the chemicals to be delivered (For example, specify the location where the chemicals are to be stored, purchase the personal protection equipment needed to transfer them, and provide the necessary labels and warning signs)

10. Do not accept chemicals delivered in damaged containers or with faulty labeling.

11. Check the updated list of hazardous chemicals to update your information on the latest chemicals delivered to the lab.

12. Use GHS standards for a uniform categorization and labeling of chemicals used in the lab.

13. Smoking is strictly prohibited in the labs.

14. All hazardous chemicals should be kept at eye level or below.

15. Use shock-proof equipment when manually handling and transferring chemicals.

16. If using a wheelbarrow to transfer chemicals, make sure it is stable enough and does not tilt to the sides to minimize chemical spillage.

17. Whenever possible, use fright elevators to carry chemicals to minimize any potential contact between the chemicals and the personnel.

18. Use a forklift or trolleys to carry and transfer gas cylinders and lab containers.

19. Do not roll over or drag gas cylinders on the ground.

20. When carrying gas cylinders, ensure the safety cap is locked.

21. It is prohibited to carry or transfer chemicals in personal vehicles.

22. If chemicals should be transferred to a new place where hot working and cold working are to be done, the process should be mentioned in work permit.

23. The requirements set in the local, national, and international regulations should be met when transferring any chemical out of the lab. Lab stuff includes chemicals, compressed gas cylinders, samples, cooling materials, and tools and equipment containing hazardous substances. These lab stuff should be sorter, packages, and marked before their transfer.

## 9. Emergency Planning

- The following information should be provided at the main entrance doors of labs where chemicals are used and stored:

- Names and emergency phone numbers of lab supervisor and all other officials who should be contacted when a fire, accident, or chemical spillage occurs.

- Certain hazards that might be present within the lab (e.g., hazards associated with laser radiation, cylinders, and radioactive substances)

Figure 12. A sample of information stickers provided at the lab’s main entrance door

- Safety instructions requiring all individuals entering the lab to wear the special personal protection equipment and identifying areas with restricted access

- Prohibitions (smoking, eating, drinking, etc.)

- NFPA Hazard diamond (based on NFPA 704 Standard) Emergency gathering areas map provided at all entrance doors Signs indicating all waste collection points

- Sings indicating where safety equipment can be found (Safety equipment include manual fire extinguishers, safety showers, eye washes, etc.)

- It is mandatory to provide labels, signs, and warning tapes around the areas where chemicals and lab equipment are stored

- The lab head or supervisor should explain the emergency response plan to the lab personnel.

- The emergency response plan should include the following details:

- Evacuation procedure: Emergency evacuation routes and gathering areas should be marked (emergency situations include accidents like leakage and spillage of chemicals).

- Warning alarm system operation: Mark the location of warning alarm switches. Different alarms can be set for different emergency situations (fires, chemical leakage, etc.).

- Emergency Shut-off system: Locations of emergency shut-off systems are marked (e.g., systems used to shut off gas discharge, compressed air exhaust and electricity).

## 10. Chemicals Release and Spillage and Relevant Cleaning Methods

- All NIOC’s subsidiaries and associate companies shall be obliged to develop their own release and spill response and cleanup procedures to be followed in emergency situations.

- Neutralization and cleanup procedures for specific substances should be identified.

- Special Spill kits should be provided and made available.

11. Equipment Repair, Maintenance, and Inspection

- All equipment should be regularly inspected for potential faulty electrical cables, corrosion, leakage, contamination, safety system failures, etc.

- Ensure all alarm, protection, and interlock systems and other safety equipment are properly functioning. A label or card recording the inspection date should be attached on the equipment (safety showers, manual fire extinguishers, lab hoods, etc.).

- Safety masks should be worn in situations where chemicals are likely to spill or form aerosols. Special safety masks should be worn when dealing with radiologic substances or laser radiation.

- Each lab should be equipped with at

- least on multi-purp

- ose (ABC) fire extinguisher. This fire extinguisher should be easily accessible. When developing emergency response plans, manual fire extinguishers should be marked white and red signs. These signs should be installed on top of the manual fire extinguishers.

- The annual inspection reports should be entered on the special inspection and recharge labels attached to the manual fire extinguishers.

- The lab personnel should have access to appropriate first aid kits. The green and white signs on top of the kits indicates their location. The contents of first aid kits should be approved by the regional health care center and be periodically controlled and recharged by the relevant authorities.

- Flammable liquid cans have a sealed lid and are fire-retardant. These cans are used to store flammable liquids with a maximum volume of 4 liters.

- Flammable liquid cans should be labeled according to the relevant standards.

- Chemical storage freezers and refrigerators should not be used to store other materials.

- Chemical storage freezers and refrigerators should not be modified.

- A label should be attached on the refrigerators and freezers specifically used to store food and drinks, refraining users from storing chemicals in such units.

- Food and drinks should not be stored in chemical storage freezers and refrigerators. A label should be attached on these freezers and refrigerators, refraining users from storing food and drinks in such units.

Figure 13. A sign prohibiting food and drinks to be stored in a chemical storage refrigerator

## 12. Personal Protection Equipment

- The minimum personal protection equipment needed for the lab personnel include:

- Lab uniform (lab coat, apron, etc.)

- Safety glasses/ goggles

- Safety gloves

- Safety shoes (closed shoes bot at the front and at the back)

- Other protection equipment, such as special goggles, mask, gloves resistant to cutting and heat, safety earmuffs, and respiratory masks

- Sandals are prohibited in the lab.

- Personnel dealing with flammable substances should wear FR (fire-resistant) uniforms.

- The lab supervisor should introduce all the individuals working in the lab to the relevant personal protection equipment training programs.

- All personal protection equipment should be periodically inspected so that potential defects be detected. The user himself/herself should also inspect such equipment before using them.

- Disposable personal protection equipment such as disposable gloves should be properly disposed after each use.

- Safety mask should be used along safety glasses and goggles in situation where your eyes and facial skin should be protected against the potential hazards, Safety mask alone does not provide eye protection.

- Safety gloves protect hands against different harmful substances, cuts, scratches, and chemical and thermal burns.

- Safety gloves should be selected according to the requirements of the situation. When selecting safety gloves, their thickness, permeability (minute), and corrosion-resistance should be taken into consideration.

- Use safety earmuff in situation where the noise level exceeds 10 dB..

- Use of respiratory equipment is mandatory especially in cases where engineered air pollution control systems installed in the lab are not adequate.

- Medical tests, congestion tests, and training is mandatory before using respiratory equipment.

- Wear safety glasses with side protection when working in la ab where there is likely that objects, regardless of their size, be thrown at your eye.

- Wearing short-sleeve shirts in the lab is strictly prohibited.

## 13. Training

The lab head and supervisor shall be obliged to make sure all the personnel receive the necessary training before starting their work in the lab and are regularly retrained.

- Safety and health training programs should be periodically organized for the lab personnel.

- The lab supervisor should evaluate the personnel’s protection skills and their understanding of safety and health issues and if required organize a retraining program. Is some cases, retraining is needed due to the changes in the working conditions.

- Documents related to the training of the lab personnel should

- be kept by the relevant department for at least one year.

- All lab personnel including the employees and interns should receive the necessary special training related to their work. These involve:

- Location and content of safety instructions available in the lab

- Physical, chemical, biological, and ergonomic hazards, risks associated with working with electrical equipment, ionizing and laser radiation hazards, warning signs and marks, and chemical exposure limits

- Places where documents on the potential hazards of chemicals and relevant safety measures are kept

- Protection measures to prevent injuries

- Emergency response procedures Methods to detect chemical, biological and radioactive substances present or released in the lab Instruction on how to use first aid kits and follow the rescue and relief procedures in emergencies

- Proper management and disposal of laboratory waste

- Proper storage of safety and health documentations

## 14. Differences between Safety Data Sheets (SDS) and Safety Labels

Safety labels associated with chemicals give use a brief introduction to the chemicals’ potential hazard. All containers holding chemicals should be labeled according to the relevant standards. Manufacturers and suppliers of chemicals are obliged to properly label their products. Even containers carrying chemicals to be readily used should at least be labeled with the name of the chemical they hold. Safety labels, just like safety data sheets (SDSs), give us information on the properties of chemicals. Thus, SDSs can be used as references when preparing safety labels.

The information provided by the labels attached on the containers should comply with the requirements set in the GHS standard, and follow the instructions given in the table below. A uniform labeling procedure should be used in all NIOC’s subsidiaries and associate companies.

A full label should contain the following information:

| Name of substance: C2H6O (alcohol) CAS NO.: 64-17-5 |  |  |
| --- | --- | --- |
| Warning word: Danger |  |  |
| Hazards: Conjunctiva irritation Skin irritation Nasal glands irritation Poisoning, should not be swallowed, eating and drinking near the substance is prohibited |  |  |
| First aid measures: Remove contact lenses from the eyes (if worn), wash the eyes with the flowing water (cold) without using your hands for about 15 minutes, refrain from using eye ointments, visiting a doctor No case was reported on skin contact Move the individual to a place where fresh air is available, use an oxygen cylinder in case of breathing difficulties, visit a doctor Refrain from attempting to vomit, loosen the buttons, oral feeding is prohibited, visit a doctor |  |  |
| Storage conditions: The substance should not be exposed to direct sunlight. It should be stored in an enclosed, cool, and ventilated place with an appropriate temperature. |  |  |
| Safety cautions: |  |  |
| Disposal method: The substance should not be disposed in flowing waters or near aquifers, as doing so will have both shore- and long-term health and environmental consequences. |  |  |
| Personal Protection Equipment: Safety glasses Uniform and safety gloves Normal respiratory mask |  |  |
| Importing/manufacturing company: |  |  |

Table 1. A sample of pre-made secondary safety labels

### 14-1 Proper Chemicals Labeling

- All chemicals used and stored in the lab should be clearly labeled so that all the lab personnel, visitors, etc. can easily become aware of the potential hazards associated with these substances.

- Labeling should be done according to the GHS system.

- Chemicals should be stored according to the storage conditions provided by the manufacturing company.

- All containers should be stored in appropriate conditions and properly labeled. The containers’ lids should be closed

- All hazardous substances should be stored in special containers. These containers should be kept in safety cabinets or certain shelves.

- Incompatible substances and chemicals which are likely to react when in close contact should not be stored in the same place (Appendix 2)

- Chemicals should be sorted and separated according to their hazard level. (e.g., flammable liquids, organic acids, and oxidants) (Appendix 1).

- All cabinets and shelves should be secured to prevent containers from falling.

- Chemicals, especially hazardous chemical liquids, should be stored at altitudes below 7 ft (213.36 cm).

- If possible, all chemicals should be stored at altitudes below 7 ft.

- A list of all chemicals annually used in the lab should be prepared.

## 15. Safe Use of Chemicals

- Carefully read the container's label containing the information provided by the manufacturing company.

- Carefully read the SDS before using the chemical to become aware of all the associated potential hazards.

- Carefully exercise all the safety cautions when using the chemicals.

- If hazardous fumes and particulates are generated during the work, the operation should strictly be conducted under a hood or an appropriate ventilation system.

- When using chemicals, containers carrying incompatible substances should not be places close to each other.

- Do not exercise mouth pipetting or use your mouth to pick up chemicals from inside the containers. Use a pro pipette or other special tools instead.

- Do not smell or taste chemicals.

- Make sure the container’s lid is firmly closed.

- The chemicals should only be taken from the storage or places where they are kept in the required amounts.

- Use the necessary personal protection equipment when dealing with chemicals.

## 16. Lab Inspection

- Inspections should be performed periodically and based on the relevant

- checklist. In the event of any malfunction in the lab, the lab supervisor should submit his/her inspection request to the HSE Department.

- Inspection documents and records shall be stored in the lab and HSE Department’s Archive for at least one year.

## 17. Special Measures for Certain Hazards

17-1 Compressed and Liquid Gases

- Ensure gas cylinders are kept away from thermal sources and their safety cap is securely locked when not in use.

- Cylinders containing hazardous gases should be stored in a ventilated cabinet or en enclosed area equipped with an appropriate ventilation system.

- Never roll over or drag gas cylinders on the ground. Use a trolley of other appropriate equipment to move around gas cylinders.

Figure 14. An appropriate equipment from moving around compressed gas cylinders

- The safety cap should be securely locked when cylinders are being moved around.

- Regulators, manifolds, and safety valves should be properly used and maintained.

- Always use safety goggles when dealing with compressed or liquid gases. The need to use other protective equipment (e.g., protective mask, insulated gloves, chemical-resistant gloves and aprons) depends on the type of the liquid or compressed gas.

- After the compresses or liquid gas system is connected, check all available connections.

- Special detectors can be used for checking these connections. Leakage inspection and control should be carried out regularly.

- In the event of a leakage, necessary safety cautions should be exercised and the required measures be taken.

17-2 Corrosive Substances

- When dealing with corrosive substances, use goggles that are resistant to chemical spillage. If necessary, wear safety masks, aprons, and boots.

- Acids and bases should be added to the water in a slow and swift manner. Never add water to acids or bases.

- Acids and bases should be sorted and stored in separate places.

- Organic and mineral acids should be separated.

- Oxidant acids (nitric acid, perchloric acid, chromic acid, and sulfuric acid) should be kept away from one another. Use appropriate neutralizing agents to clean and wash up spilled corrosive substances. These agents should be available at specific locations.

- Calcium gluconate gel should be available at places where hydrofluoric acid is used or stored. The gel is used to neutralize acids spilled on the skin. It should be also available at the clinic in appropriate amount.

- Polyethylene glycol (PEG300) should be available at places where phenol is used or stored. IT is used to neutralize phenol when spilled on skin. It should be also available at the clinic in appropriate amount.

- Perchloric acid can form explosive perchloric acid salts. The acid should always be handles under a hood.

17-3 Explosive Substances

- Explosive substances are found in two forms. The first group consists of substances produced to create explosion. The second group includes substances that become explosive by dehydration, contamination, or passage of time. TNT and picric acid are examples of explosive substances.

- Storage conditions of explosive substances should be registered in a special form.

- Explosive substances should be stored in cabinets or special storage rooms and away from other chemicals.

- Explosive substances should be storage under stable conditions.

- Explosive substances should be protected against heat, friction, and static electricity.

17-4 Flammable and Combustible Substances

- Flammable substances have a maximum flash point of 37.8°C at atmospheric level. Combustible substances have a flash point higher than 37.8°C at atmospheric level.

- Sources of spark (e.g., hot surfaces, electric equipment, and static electricity) should be kept away from flammable and combustible substances.

- Flammable liquids with volumes more than 4 liters should be stored in safety cans.

- If the total volume of flammable substances available in the lab is greater than 40 liters, they should be stored in flammable storage cabinets.

- Flammable and combustible substances should be kept away from oxidants and strong acids. Flammable liquids required to be kept in cold conditions should be stored in refrigerators specially designed for their storage. Modification of normal refrigerators to be used for the storage f flammable substances is prohibited.

- Before decanting flammable liquids, grounding and bonding of the new container should be properly carried out.

17-5 Mercury

- Mercury, even in small quantities, is considered a toxic substance and if spilled, can have significant consequences.

- Equipment containing mercury should be kept in special containers to minimize the risk of mercury spillage.

- Equipment containing substances other than mercury are recommended. For instance, ethanol-filled and electronic can replace mercury thermometers.

- Small quantities of mercury (mercury spill resulting from the breaking of a mercury thermometer) can be cleaned up by the user. However, in cases where the spillage is large, the cleanup operation should be arranged and supervised by HSE authorities.

17-6 Organic Peroxides

- Organic peroxides might react to organic substances, leading to fires or explosions. These substances are highly explosive and sensitive to heat, friction, and oxidants.

- If contaminated or improperly stored, they will automatically trigger reactions. Passage of time can also lead to their automatic reaction.

- Peroxides commonly used in lab include benzoyl peroxide, butyl peroxide, and lauroyl peroxide.

- The receipt date should be written on the container and disposal date be determined according to the relevant list.

- Organic peroxides have an expiry date of no more than 6 months from the time their container has been opened.

- Organic peroxides should be kept away from chemicals and organic materials like paper and wood.

- Organic peroxides should be stored and used according to the instruction provided by the manufacturing company.

- Organic peroxides should not be concentrated (distillation or crystallization).

17-7 Oxidants

- Oxidants’ reaction to organic substances can lead to fires or combustion. Perchloric acid and nitric acid are among the most common oxidants used in labs.

- Oxidants should be kept in separate containers away from other substances.

17-8 Peroxide-Producing Chemicals

- Peroxide-producing chemicals are substances that produce peroxides when reacting to oxygen.

- Shock, heat, or friction can cause the produced peroxide to explode. Ethyl ether, tetrahydrofuran, isopropyl ether, and potassium metal are among the peroxide-producing chemicals.

- A warning label should be attached to the containers carrying peroxide-producing chemicals. The receipt date, the date when the container was opened, and the date when it was tested should be indicated on the container.

- Peroxide-producing chemicals should be disposed of before their expiry date.

- Deformed containers or containers made of crystals should be carefully handled.

17-9 Pyrophoric Substance

- Pyrophoric substances (e.g., butyl lithium, methyl lithium, and white phosphorus) automatically burn when exposed to air. Small quantities of such substances can cause large and widespread fires.
Pyrophoric substances should be used and stored in a neutral environment.

- Pyrophoric substances should not be used in close proximity to flammable solvents.

- Before their disposal, the containers carrying pyrophoric substances should be neutralized based on the standard methods used in the lab.

- Before working with pyrophoric substances, make sure appropriate fire extinguishers are available in the lab.

17-10 Water Reagents

- Many water reagents can automatically produce heat and get on fire or give off toxic or flammable gases. Water reagents include substances like sodium, potassium, and chlorine acids, and extremely small powders of metals such as zinc.

- Water reagents should be kept away from water and humidity.

- Water reagents should be stored in a properly ventilated space.

- Water reagents should be stored in a dry and neutral environment.

- The remaining water reagents should be neutralized and disposed of according to standard lab procedures.

- Before working with water reagents, make sure appropriate fire extinguishers are available in the lab.

## 18. Physical Damage Prevention

- Lab corridors and exits should be clean and unobstructed.

- Make sure all rotatory and mobile equipment used in the lab has the required protective guards.

- Ensure ladders and stools are placed at safe positions.

- Make sure heavy objects are placed at heights less than 5 feet to minimize their risk of falling.

- Ensure adequate lighting is provided for different activities. Lighting should not cause glare.

## 19. Waste Management in Labs

All NIOC’s subsidiaries and associate companies should follow the following guidelines and instruction for waste management in their labs:

- NIOC ‘s Waste Management Guideline - NIOC-HSE-EN-GU-006-00

- HSE Department’s Waste Management Guideline - MOP-HSED-GI-301(0)

- HSE Department’s Waste Categorization and Coding Guideline - MOP-HSED-GI-302(0)

- HSE Department’s Waste Management Reporting Guideline - MOP-HSED-GI-303(0)

- HSE Department’s Landfill Construction Guideline - MOP-HSED-GI-309(0)

- HSE Department’s Landfill Monitoring Guideline - MOP-HSED-GI-310(0)

## 20. References

- UOP 45

- NFPA 70

- NFPA 231

- NFPA 30

- OSHA 3404 Laboratory-Safety-Guidance

- OSHA Factsheet-Laboratory-Safety-Chemical-Hygiene-Plan

- ANSI Z9.5 Laboratory Ventilation

- ANSI 49 Electrical Installation

- CRC - Handbook of Laboratory Safety

## NIOC ‘s Waste Management Guideline - NIOC-HSE-EN-GU-006-00

- HSE Department’s Waste Management Guideline - MOP-HSED-GI-301(0)

## HSE Department’s Waste Categorization and Coding Guideline - MOP-HSED-GI-302(0)

- HSE Department’s Waste Management Reporting Guideline - MOP-HSED-GI-303(0)

- HSE Department’s Landfill Construction Guideline - MOP-HSED-GI-309(0)

- HSE Department’s Waste Management Guideline - MOP-HSED-GI-301(0)

- Chemical Storage Conditions and Features- Ministry of Health and Medical Education - Health Department - Center of Occupational and Environmental Health

- SDS for Chemicals provided by International Labor Organization (ILO)

- Hazardous Flammable Safety Codes of Practice - Ministry of Labor and Social Affairs - Workplace Safety and Health

- Hazardous Chemicals Comprehensive Instruction and Guideline - Ministry of Health and Medical Education - Health Department - Center of Occupational and Environmental Health

## Appendix 1 - Danger Classification

Classification Of Substance

Appendix 2 - Chemical Incompatibility Table

| Number | Chemical | Incompatible with … |
| --- | --- | --- |
| 1 | Acetic acid CH3COOH | Oxidants such as chromic acid, nitric acid, heterocyclic compounds, Ethylene glycol, Perchloric acid, peroxides, permanganates |
| 2 | Acetylene C2H2 | Chlorine, bromine, copper, silver fluoride, mercury |
| 3 | Alkali metals and alkaline earth metals such magnesium powder, calcium, lithium, sodium, potassium, aluminum powder | Water, carbon tetrachloride, other hydrocarbons with chlorine, carbon dioxide, halogens |
| 4 | Anhydrous ammonia NH3 | Mercury, chlorine, calcium hypochlorite, iodine, bromine, acid Hydrofluoric |
| 5 | Ammonium nitrate NH4NO3 | Acids, metal powders, flammable solutions, chlorates, nitrites, sulfur, micro organic compounds |
| 6 | Aniline C6H5NH2 | Nitric acid, hydrogen peroxide |
| 7 | Substances containing arsenic As | Reducers |
| 8 | Azides N3- | Acids |
| 9 | Bromine Br | Ammonia, Acetylene, butadiene, butane, methane, protane (or other gases produced by petroleum), hydrogen, benzene-sodium chloride, metal powders, turpentine |
| 10 | Calcium oxide CaO | Water |
| 11 | Activated carbon C | Calcium hypochlorite, other oxidants |
| 12 | Chlorates ClO4- | Sal ammoniac, acids, metal powders, sulfur, micro organic compounds, flammable substances |
| 13 | Chlorine Cl2 | Ammonia, Acetylene, butadiene, butane, methane, protane (or other gases produced by petroleum), hydrogen, benzene-sodium chloride, metal powders, turpentine |
| 14 | Chlorine dioxide ClO2 | Ammonia, methane, phosphine, hydrogen sulfide |
| 15 | Chromic acid H2CrO4 | Acetic acid, naphthalene, camphor, glycerol, alcohol, flammable solutions |
| 16 | Copper Cu | Acetylene, hydrogen peroxide |
| 17 | Cyanides CN- | Acids |
| 18 | Flammable solutions | Ammonium nitrate, chromic acid, nitric acid, halogens, hydrogen peroxide, sodium peroxide |
| 19 | Hydrocarbons CnH(2n+2) | Fluorine, chlorine, bromine, chromic acid, sodium peroxide, other oxidants |
| 20 | Hydrocyanic acid HCN | Alkali |
| 21 | Hydrofluoric acid HF | Potassium permanganate, sulfuric acid |
| 22 | Hydrogen sulfide H2S | Metal oxides, copper powder, oxidants |
| 23 | Hypochlorites ClO- | Acids, activated carbon, ammonia |
| 24 | Iodine I2 | Acetylene, ammonia, hydrogen |
| 25 | Mercury Hg | Acetylene, folinic acid, ammonia |
| 26 | Nitrates NO3- | Metal and non-metal powders, metal sulfides, combustible solutions |
| 27 | Nitroparaffins CnH(2n+2)NO2 (n=20-40) | Acids, bases, amines, halides |
| 28 | Oxalic acid C2H2O4 | Silver, chlorites, urea |
| 29 | Oxygen O2 | Oils, grease, hydrogen, other reducers including gases, solutions, and flammable substances |
| 30 | Perchlorates ClO4- | Sal ammoniac, acids, metal powders, sulfur, micro organic compounds, flammable substances |
| 31 | Perchloric acid HClO4 | Reducers including acetic anhydride, bismuth and associated alloys, alcohol, paper, wool, grease, oils |
| 32 | Phosphor P | Air, oxygen, alkali, halogens, halogen oxides, oxidants |
| 33 | Potassium K | Carbon tetrachloride, carbon dioxide, water |
| 34 | Potassium permanganate KMNO4 | Glycerol, ethylene glycol, benzaldehyde, other reducers, sulfuric acid |
| 35 | Sodium Na | Carbon tetrachloride, Carbon dioxide, water |
| 36 | Sodium peroxide Na2O2 | Ethanol, methanol, acetic acid (glacial), acetic anhydride, benzaldehyde, carbon disulfide, glycerin, glycol ethylene, ethyl acetate, methyl acetate, furfural |
| 37 | Sulfides S2- | Acids |
| 38 | Sulfuric acid H2SO4 | Permanganate, water, aqueous solutions, reducers, chlorates, perchlorates, nitric acid |
| 39 | Water | Alkaline metals, alkaline metal hydrides, alkaline metal amides, metal alkyls including lithium alkyl, aluminum alkyl, Grignard reagents, non-metal halides including S2Cl2, SiCl4, PCI5, PCI3, BF3, BCl3, non-organic halide acids including POCl3, SO2Cl2, SOCl2, phosphorus pentoxide, calcium carbide, organic halide acids, anhydrides with lo molecular weight |
| 40 | Acetaldehyde CH3CHO | When reacting with water, it produces explosive peroxides. |
| 41 | Acetic anhydride (CH3CO)2O | Reacting with acids or alkali, it can produce polymer in the presence of rare metals. It is a strong reducer, It violently reacts with oxidants. It can react with different organic substances, halogens, sulfuric acid, and amines. |
| 42 | Acetone CH3COCH3 | It violently reacts with oxidants (chromic acid and nitric acid) and chloroform in the presence of alkali. It is incompatible with thick nitric acid. |
| 43 | Acetonitrile CH3CN | It reacts with aqueous acids and bases and generates toxic fumes. It reacts with strong oxidants. It attacks some kinds of plastics, rubbers, and coatings. It breaks through burning, creating hydrogen cyanide and nitrogen oxide. |
| 44 | Benzene C6H6 | Containers should be kept in a well-ventilated place away from sources of spark. It should be handled in well0ventilated hoods. Use safety goggles and nitrile or PVC gloves when handling benzene. Earthing would prevent electric charge. |
| 45 | Carbon tetrachloride CCl4 | It will deteriorate and generate corrosive fumes and gases (hydrogen chloride and phosgene) when in contact with hot surfaces or flames. Sometimes, it reacts with metals such as aluminum, magnesium, and zinc. |
| 46 | Chloroform CHCl3 | Strong bases, metals such as aluminum, magnesium zinc powder, strong oxidants |
| 47 | Carbon dioxide (sold, ice, dry) CO2 | Alkaline metals, strong bases |
| 48 | Acrolein CH2=CHCHO | Oxidants, acids, alkali, ammonia, amines. It is easily polymerized unless it is harnessed by hydrocarbons. It can generate peroxides sensitive to long-term impacts. |
| 49 | Cyanogen bromide BrCN | It deteriorates when exposed to heat. When in contact with acids, it creates hydrogen cyanide which is highly toxic and flammable and hydrogen bromide which is corrosive. It reacts with strong oxidants. |
| 50 | Acetonitrile CH3CN | It reacts with aqueous acids and bases and generates toxic fumes. It reacts with strong oxidants. It attacks some kinds of plastic, rubber, and coating. It breaks through burning, creating hydrogen cyanide and nitrogen oxide. |
| 51 | Diethyl ether C2H5OC2H5 | When exposed to air or light, it can generate explosive peroxides. It can mix with oxidants and halogens. |
| 52 | Dimethylamine (CH3)2NH | It can react with oxidants, mercury. |
| 53 | Dinitrophenyl hydrazine NHNH2 Hydrazino Dinitrobenzene C6H3(NO2)2 | It violently reacts with reducers and oxidants. |
| 54 | Dioxane C4H6O2 Diethylene dioxide | It can generate explosive peroxides. It can violently react with strong oxidants and thick acids. It will explode when in contact with some catalyzers. It attacks many plastics. |
| 55 | Ethanol CH3CH2OH | It reacts with strong oxidants. |
| 56 | Ethanolamine H2NCH2CH2OH Ethanolamine | It reacts with strong oxidants. |
| 57 | Methanol CH3OH | It violently reacts with oxidants. It can also react violently with magnesium and bromine. It explodes when in contact with chloroform and sodium. |
| 58 | Nitric acid HNO3 (50-70%) | Acetic acid, chromic acid, hydrocyanic acid, aniline, carbon, hydrogen sulfide, bases, metals, etc. |
| 59 | Nitrobenzene C6H5NO2 | When burnt, it generates corrosive gases. Like nitrogen oxides, it can violently react with strong oxidants and reducers, contributing to fires and explosions. It attacks many plastics. It generates explosive substances (thermally unstable thermal) when in contact with a wide range of organic and non-organic substances, |
| 60 | Perchloric acid HClO4 | Flammable substances and reducers: acetic anhydride, bismuth and its alloys, alcohol, metals, paper, organic substances |
| 61 | Phosphorus pentoxide P2O5 | It aqeous solution is a strong acid. It violently reacts with bases and is corrosive. It violently reacts with perchloric acid and contributes to fires and explosions. It violently reacts with water, generating acids. |
| 62 | Potassium hydroxide KOH | It violently reacts with acids, nitrobenzene, and many detergents. It combines with water, producing heat. It should be stored in containers with securely fastened lids. |
| 63 | Silver Ag | It violently reacts with strong oxidants and acids. It reacts with water at 50 degrees centigrade, generating flammable hydrogen and selenium acids. When exposed to phosphor and metals such as nickel, potassium, platinum, sodium, and zinc along with mild heat, it generates light. |
| 64 | Silver nitrate AgNO3 | Ammonia solutions generate explosive silver nitrate sediments in the presence of bases or glucose. It reacts with methanol, generating explosive substances. It reacts with acrylonitrile, generating explosive polymers. It can contribute to explosion when combined with charcoal, phosphor, magnesium, or sulfur. |
| 65 | Sodium azide NaN3 | It reacts with bromine, carbon disulfide or chrome chloride, creating explosions. In its solid form, it reacts with heavy metals such as copper, lead, and mercury, generating explosive azide salts. When exposed to acid, it generates highly toxic and explosive gases. |
| 66 | Sodium cyanide NaCN | When exposed to acid or water containing carbon dioxide, it can generate hydrogen cyanide which is a highly toxic gas. It can react with nitrites generating explosive mixtures. |
| 67 | Tehtrahydrofuran C4H8-O Dimethylene oxide Tetramethylene Oxide | It can react with strong oxidants and bases and some metal halides, contributing to fires and explosions. Tehtrahydrofuran can be polymerized in the presence of cationic initiators. |

No. 17, Yaghma Alley, Jomhouri-e Eslami St. (btwn Hafez St. & Valiasr Ave.), Tehran, Tehran, Iran

Phone: +98 21 66726016, Fax: +98 21 6609723, SMS System: +98 10007615

HSEE@NIOC.IR HTTP://HSE.NIOC.IR